INSTRUCTION MANUAL FOR INCUBATOR DSI-D & DSI-DF SERIES
[image: image1.png]


[image: image15.jpg]



[image: image16.png]



[image: image17.png](&)




[image: image18.png]



Warnings:
* Make sure that the door is closed well to prevent heat loss and power drain.

* The surface is hot. Do not touch.

* Do not put any objects on the top of the instrument.

* Please keep the environment ventilated.
1. Connect with power supply. Please make sure the power voltage is the same as the voltage shown on the label stuck on the back of machine.

2. Put your sample in and close the door.

3. Press “ON/OFF Button”, then you will see the “Power indicator’ becomes green. Wait for about 5 seconds; you can see that on the display screen, the top row is the present temperature value (PV) and the bottom row is the setting temperature value (SV).

[image: image19.png]



[image: image20.png]


[image: image21.png]


[image: image22.png]


4. Set temperature: you will see           . Press    or    to adjust SV value and then press    to enter the value.
[image: image23.png]ECME



[image: image24.png]ENE



5. Set time: Select time mode: Continouos mode     or Timer mode    . 
[image: image25.png](PV)
Sil(SV)



5.1 Press     to set Continuous mode. OUT2 indicator on display screen will disappear. The setting is finished. Go to Step 6.
[image: image26.png]HI5(PV)
gggy (remaining




[image: image27.png]LAB. INCUBATOR

Power

L, 1 = indicator
' '.b‘i:

OUT1 REMOTE - 7
G’} ouT2 EV

l
Y AN

Timer mode key Y Continous mode key
_J ‘
ON/OFF

S« v o

Select key Down key Up key

DIGISYSTEM Laboratory Instruments Inc.



5.2 Press    to set Timer mode. OUT2 indicator on screen will appear.

[image: image28.png]dEI/

d5¢





   5.2.1 You will see       or       . 

5.2.2 Use  [image: image29.png]SSR

’_H_H_‘:m” 2 1

Froon

CEES

DTA330VA-0
>0 0
L Lo Ls Ns

ON

OFF




  or  [image: image2.png]


  key to select time unit [image: image3.png]


 (hour) or [image: image4.png]


 (minute). Press [image: image5.png]


 to enter the value. Press  [image: image6.png]


  to next setting.
5.2.3 You will see       . Use    or    key to adjust the time value. Press     to enter the value. Press    to next setting.


5.2.4 You will see            . Press    and you will see           . The setting is finished.
Remark: If you want to interrupt the timer mode, you can press     to be continuous mode.

6. After finishing the work, please press “ON/OFF Button” to turn off the power, and then the “Power indicator’ becomes red.
Remark:
1. When the “OUT1” indicator on the display screen is sparkling, it means that the instrument is heating.

2. After the PV value reaches to SV value and tends to be stable, the “ALM1” indicator lamp will appear if the PV value exceeds the SV value by setting point (10℃). In this situation, please turn the power off, open the door and wait for about 30 minutes, and then restart it.

3. If you feel the PV value is much higher or lower than the SV value, you can adjust the “Air vent” on the back of machine to balance the temperature. That is, if the PV value is too high, you can open the “Air vent”; if the PV value is too low, you can close the “Air vent”.
Maintenance and care:
Keep the machine clean all the time, so no routine maintenance is required. Cleaning can be done with a damp cloth. Avoid the use of solvents as they may damage the product housing and inner chamber.
Circuit Diagram:


Other functions:

Note: Before executing other functions , please make sure that the continuous mode is selected. 

     If you see OUT2 indicator on the screen, please press     to set Continouos mode.
1. Other functions include PV shift (offset) value, setting alarm point, setting temperature unit, auto tuning function.
	Action
	Display

(Initial value)
	Explanation
	Adjustment

	Press    
	[image: image7.png]



(0)
	PV shift (offset) value. If PV value is not correct, you can use this function to calibrate by your own thermometer.
	Press   or    to adjust the value. Then press    to enter the value.

If you want to go back to PV/SV display, press    again.

	Press 
	[image: image8.png]



(10)
	Upper-limit point for alarm. When the difference between PV and SV is over upper-limit, the alarm indicator (ALM1) on display screen will be flashing.
	

	Press 
	[image: image9.png]



(10)
	Lower-limit point for alarm. When the difference between SV and PV is over upper-limit, the alarm indicator (ALM1) on display screen will be flashing.
	

	Press
	[image: image10.png]



(C)
	Temperature unit. C or F.
	

	Press
	[image: image11.png]



(OFF)
	Auto-tuning function. When auto-tuning function is on, the “AT” indicator” on display screen is flashing. When the function is off, the light of the “AT” indicator” is extinguish.
	


2. Restore default value: 

2.1 Press      more than 3 seconds, and then you will see        。 

   2.2 Press     and then press     , you will see       is flashing. Please press    to enter the value.
      Then the default value is restored.
Service Manual
Troubleshooting Guide
Caution: Always disconnect the power cord before troubleshooting.
	Trouble
	Cause
	Remedy

	Instrument inoperative
	Power cord not connected to outlet. 
	Plug instrument in.

	
	Dead power output.
	Change to different output.

	
	No fuse Breaker is off
	Press the breaker back of the machine, and check if the current is overload.

	
	Electronic element broken
	Contact your distributor for repair.

	
	Power switch is broken
	Contact your distributor for repair.

	Controller can’t control the temperature
	Sensor is broken
	Contact your distributor for repair.

	
	Controller
	Contact your distributor for repair.

	
	Heater is broken
	Contact your distributor for repair.

	Temperature is not stable
	The initial setting value is not suitable for the ambient air temperature where the machine is located
	Refer to the other functions, execute the Auto-tuning function.

	PV value is not correct
	User’s calibration thermometer is different from the factory’s calibration thermometer
	Refer to the other functions, execute the function for PV shift (offset) value.


[image: image12.png]


[image: image13.png]


[image: image14.png]



Air vent





Controller








DIGISYSTEM LABORATORY INSTRUMENTS INC.               www.digisystemlab.com                        REV.1-2017

