# INSTRUCTION MANUAL FOR ROLLER MIXER

Power indicator

ON/OFF Switch

Power socket

Rollers

Warnings:

* Do not remove the mixer during operation.

* Place the mixer securely away from the edge of bench to avoid the mixer falling off the bench top.

* Do not disassemble the mixer by yourself. In the event of a malfunction, the mixer should be repaired by a qualified technician. Before repairing, turn off the power.

Steps:

- Place the mixer on a flat and stable bench top. Connect cable with power socket and a properly grounded outlet which corresponds with the electrical requirements on the label back on machine.

- Put the specimen into a container and tight closed , and then place the container between ‘Rollers’.

- Place ‘ ON/OFF Switch’ at ‘ON’ position, and then the power indicator is illuminated. Then the machine starts rolling.

- Containers can be loaded and unloaded during the machine rolling.

Maintenance and care:

Keep the machine clean all the time, so no routine maintenance is required. Cleaning can be done with a damp cloth. Avoid the use of solvents as they may damage the machine housing.

Circuit Diagram

Troubleshooting Guide

Caution: Always disconnect the power cord before troubleshooting.

| Trouble | Cause | Remedy |
| --- | --- | --- |
| Instrument inoperative | Power cord not connected to outlet. | Plug in the power cord. |
|  | Dead power output. | Change to different output. |
|  | Fuse blown. | Change a new fuse. |
|  | Defective electronic component. | Contact your distributor for repair. |
| Rollers doesn’t rotate | Switch is at “OFF” position. | Place the switch at “ON” position. |
|  | PC board is broken. | Contact your distributor for repair. |
|  | Motor is broken. | Contact your distributor for repair. |
|  | Switch is broken. | Contact your distributor for repair. |
|  | Drive pully oiled on the rubber ring. | Contract your distributor for cleaning up the oil. |
|  | Rubber ring is broken. | Contact your distributor for repair. |
|  | Plastic fixer for driven pulley is broken. | Contact your distributor for repair. |
|  | Plastic gear is broken. | Contact your distributor for repair. |
|  | The motor, plastic gears, and rollers don’t joint well. | Contact your distributor for repair. |