- INSTRUCTION MANUAL FOR DIGITAL ROLLER MIXER

Rollers

ON/OFF Switch

Control Panel

Power socket

Warnings:

* Do not remove the mixer during operation.

* Place the mixer securely away from the edge of bench to avoid the mixer falling off the bench top.

* Do not disassemble the mixer by yourself. In the event of a malfunction, the mixer should be repaired by a qualified technician. Before repairing, turn off the power.

Steps:

1.	Place the roller mixer on a flat surface. Put the specimen into a container and tight closed , and then place the container between ‘Rollers’.

2. Turn on the ‘Power Switch’ behind the machine. First, the SPEED and TIME display will show the speed and time value that you set last time. If the previous timed mode setting is Cont. mode (Continuous mode), you will see ‘cont’ on TIME display.

- Setting time: Press the      key to select the time mode (Timed mode or Continuous mode).

  - If the continuous mode is selected, the TIME display will show ‘cont’.

  - If the timer mode is selected, the previous setting value will show on the TIME display.

    - Press     key, and the digits on the TIME display is flashing. (Time display is Hour: Minute; for the last one minute, it will display the remaining seconds)

    - Press     key or     key to adjust the value.

    - Press     key or any other keys to save the time value.

- Setting speed:

  - Press     key, and the digits on the SPEED display is flashing.

  - Press     key or     key to adjust the value.

  - Press     key or any other keys to save the speed value.

- Press     key to start the machine. During the operation, you can press     key to stop the machine anytime.

- Containers can be loaded and unloaded during the machine rolling.

Maintenance and care:

Keep the machine clean all the time, so no routine maintenance is required. Cleaning can be done with a damp cloth. Avoid the use of solvents as they may damage the machine housing.

Circuit Diagram:

Service Manual

# Troubleshooting Guide

Caution: Always disconnect the power cord before troubleshooting.

| Trouble | Cause | Remedy |
| --- | --- | --- |
| Instrument inoperative | Power cord not connected to outlet. | Plug in the power cord. |
|  | Dead power output. | Change to different output. |
|  | Fuse blown. | Change a new fuse. |
|  | Defective electronic component. | Contact your distributor for repair. |
| Error message: run | Before the motor stops completely, turn off the ‘Power Switch’ behind the machine. | Turn on the ‘Power Switch’ behind the machine, after the motor stops completely. |
| Rollers don’t rotate | Power is not on. | Turn on the ‘Power Switch’ behind the machine. |
|  | Start/Stop key is not pressed. | Press the Start/Stop key. |
| Rollers don’t rotate (Error message: Er-P) | PC board is broken. . | Contact your distributor for repair. |
|  | Motor is broken. | Contact your distributor for repair. |
|  | Drive pully oiled on the rubber ring. | Contract your distributor for cleaning up the oil. |
|  | Rubber ring is broken. | Contact your distributor for repair. |
|  | Plastic fixer for driven pulley is broken. | Contact your distributor for repair. |
|  | Plastic gear is broken. | Contact your distributor for repair. |
|  | The motor, plastic gears, and rollers don’t joint well. | Contact your distributor for repair. |
| Rollers keep rotating when time is up | PC board is broken. | Contact your distributor for repair. |
| Speed is abnormal | PC board is broken. | Contact your distributor for repair. |
| No signal on LED DISPLAY | PC board is broken. | Contact your distributor for repair. |

Remark: To clear the error message, please turn off the power.