# INSTRUCTION MANUAL FOR VORTEX MIXER VM-2000 & VM-2000-C

Head

Speed controller

Switch

(AUTO/OFF/TOUCH)

VM-2000-C                   VM-2000

Warnings:

* Do not remove the mixer during operation.

* Place the mixer securely away from the edge of bench to avoid the mixer falling off the bench top.

* Do not disassemble the mixer by yourself. In the event of a malfunction, the mixer should be repaired by a qualified technician. Before repairing, turn off the power.

* Clean up the head once the user finds any liquid on the head, otherwise the head will get damaged easily.

Steps:

- Place the mixer on a flat and stable bench top. Plug the mixer into a properly grounded outlet which corresponds with the electrical requirements on the label back on machine.

- Place the ‘Switch’ at ‘TOUCH’ position. The unit will run as soon as touch on the ‘Head’. The ‘Power indicator’ lights.

- Place the ‘Switch’ at the ‘AUTO’ position, the unit will run automatically.

- Turn the ‘Speed controller’ to the required speed.

Warning: Do not make the motor race rotation to avoid damage. After finishing operation, please place the switch at ‘TOUCH’ or ‘OFF’ position.

Maintenance and care:

Keep the machine clean all the time, so no routine maintenance is required. Cleaning can be done with a damp cloth. Avoid the use of solvents as they may damage the machine housing.

Circuit Diagram

Service Manual

Troubleshooting Guide

Caution: Always disconnect the power cord before troubleshooting.

| Trouble | Cause | Remedy |
| --- | --- | --- |
| Instrument inoperative | Power cord not connected to outlet. | Plug in the power cord. |
|  | Dead power output. | Change to different output. |
|  | Defective electronic component. | Contact your distributor for repair. |
| Head doesn’t rotate | Switch isn’t at ‘AUTO or “TOUCH’ position. | Place the switch at ‘AUTO or “TOUCH’ position. |
|  | Eccentric brass rod doesn’t work well. | Contact your distributor for repair. |
|  | Switch is broken. | Contact your distributor for repair. |
|  | Motor is broken. | Contact your distributor for repair. |
|  | PC board is broken. | Contact your distributor for repair. |
| Speed can’t not be controlled | PC board is broken. | Contact your distributor for repair. |
|  | Speed controller is broken. | Contact your distributor for repair. |
| There is a big noise during operation | Eccentric brass rod doesn’t work well. | Contact your distributor for repair. |
|  | Belt which is to fix the eccentric brass rod doesn’t work well. | Contact your distributor for repair. |
| The “TOUCH” function can’t work | Switch isn’t at “TOUCH’ position. | Place the Switch at ‘TOUCH’ position. |
|  | The starting button under the motor is broken. | Contact your distributor for repair. |