# INSTRUCTION MANUAL FOR SUSPENSION ROTATOR SR-100

Dimple mat

ON/OFF Switch

Warnings:

* Do not remove the rotator during operation.

* Place the mixer securely away from the edge of bench to avoid the rotator falling off the bench top.

* Do not disassemble the rotator by yourself. In the event of a malfunction, the rotator should be repaired by a qualified technician. Before repairing, turn off the power.

Steps:

- Place the rotator on a flat and stable bench top. Plug it into a properly grounded outlet which corresponds with the electrical requirements on the label back on machine.

2.  Put the specimen into a container and tight closed , then place the container on the ‘Dimple mat’.

3.  Place ‘ON/OFF Switch’ at ‘ON’ position, and then the machine starts rocking.

4.  Containers can be loaded and unloaded during the machine operating.

Maintenance and care:

Keep the machine clean all the time, so no routine maintenance is required. Cleaning can be done with a damp cloth. Avoid the use of solvents as they may damage the machine housing.

Circuit Diagram

Troubleshooting Guide

Caution: Always disconnect the power cord before troubleshooting.

| Trouble | Cause | Remedy |
| --- | --- | --- |
| Instrument inoperative. | Power cord not connected to outlet. | Plug in the power cord. |
|  | Dead power output. | Change to different output. |
|  | Defective electronic component. | Contact your distributor for repair. |
| Power is on, but the power switch is not illuminated. | Fuse is burnt. | Replace a new fuse. |
|  | ON/ OFF switch is broken. | Contact your distributor for repair. |
| Motor doesn’t work. | Motor is broken. | Contact your distributor for repair. |
|  | ON/ OFF switch is broken. | Contact your distributor for repair. |