# INSTRUCTION MANUAL FOR DSB-500 & DSB-1000

Names of parts:

(1) Controller

(3) Thermometer

(5) Power Switch

(7) Temperature Control

(9) Fine Adjuster

(2) Tank

(4) Opening for Thermometer

(6) Pump Circulator

(8) Heater Lamp

(10) Heater

1. Take “Controller” (1) out and set it into “Tank” (2) as Figure 1.

2. Pour water into “Tank” until water level covers half height of “Tank”.

3. Insert “Thermometer” (3) into the “ Opening for Thermometer” (4), and to be sure that the end of “Thermometer” is under the water level. (In order to insert “Thermometer” into the “Opening for Thermometer” easily, you can wet “ Thermometer”.)

4. Before connecting the power supply, please check if “Power Switch” (5) is placed at “OFF” position.

5. Place “ Power Switch” at “ON” position, then switch lamp lights up and “Pump Circulator” (6) starts running.

6. Adjust “ Temperature Control” (7) to the required Temperature scale, and then “ Heater Lamp” (8) lights up. When water temperature reaches the required temperature scale, “ Heater Lamp” becomes off.

7. Check the “ Thermometer” if the temperature is the same as the required temperature. If the thermometer temperature is under the required temperature, turn “ Fine Adjuster” (9) clockwise slightly; if the thermometer temperature is over the required temperature, turn “ Fine Adjuster” (9) counterclockwise slightly. Check and adjust until the thermometer temperature is the same as the required temperature.

## Warning:

If water level is under the bottom of “Heater” (10), you must pour water into “Tank” to cover the heater.

SERVICE MANUAL

Caution: Always disconnect the power cord before troubleshooting.

| Trouble | Cause | Remedy |
| --- | --- | --- |
| Instrument inoperative | Power cord not connected to outlet. | Plug instrument in.. |
|  | Dead power output. | Change to different output. |
| Controller can’t control the water temperature | Sensor is broken | Replacing a sensor |
|  | Temperature control knob | Replacing a control knob |
|  | Heater is broken | Replacing a heater |
| Pump Circulator doesn’t work well | Pump Circulator is broken | Replacing a Pump Circulator |
|  | Foreign bodies blockaged the pump Circulator | Sweep away the foreign bodies |

### CIRCUIT DIAGRAM

# Text boxes

Figure 2

Figure 1

(10)

(4)

(2)

(5)

(9)

(8)

(1)

(7)

(3)

(6)